REPORT: COST-STSM-TD1002-181112-024118

Evaluation of nanomechanical properties of cells and soft gels using Atomic Force Microscopy

Georgios PA Michanetzis1, Tomas Luque2 and Daniel Navajas2

1 STSM Applicant; Laboratory of Biomechanics and Biomedical Engineering, University of Patras, Greece

2 Host; University of Barcelona, Spain

Period: 2012-11-18 to 2012-11-24

Abstract

Atomic force microscopy (AFM) has been used over the last fifteen years in cell biology, not only as a microscopy method acquiring high resolution images, but also as a useful “nanoindentation” tool for the evaluation of the viscoelastic properties of living cells.  However, many times the results of these studies are quite doubtful as researchers often use different and not well characterised methodologies and do not provide adequate information about data acquisition, processing and evaluation for the different AFM systems used.

The scope of this COST-STSM was the evaluation of surface nanomechanical properties of cells and soft gels using AFM, emphasizing on the methodology and the problems of the technique. As a first outcome of this work the urgent need for standardisation of the methods used was recorded.

Introduction

For more than fifteen years Atomic force microscopy (AFM) has been proved a useful tool for imaging in the case of cell biology, as well as in other sciences, when high resolution images are required. However, the possibility to use AFM also for the quantification of normal and lateral forces exerted on the AFM tip, while interacting with the surface of materials, has turned the AFM to a very useful nanoindentation tool to evaluate surface nanomechanical properties of very soft materials, where classic nanoindentation tips (fe Berkovich) fail to provide meaningful results.

AFM measurements have been used in numerous - more specific (1-7) or in a review form (8,9) - studies, to evaluate cell mechanics in the case of normal, diseased, aged cells etc aiming in a better understanding of cell mechanotransduction and of course, eventually, in the establishment of novel diagnostic techniques. However, the main problem of these studies is that they very rarely explain in details the methodology followed, assumptions and limitations in processing and evaluating the data and this, in many cases, make their results difficult to compare with results of similar studies where fe a different AFM system is used or a different software is used for the evaluation of the results. The problem appeared from the very beginning of this type of research (10) and is still present after one and a half decade of studies (11).

Our aim was to perform a sample study using human neutrophils as well as soft gels, as a reference material, to try (together with other studies) to provide standardization instructions for suggested methodology and problem solution.

Materials and Methods

Two AFM systems were used in this study: One home-made system (University of Barcelona, Spain) (Fig. 1) and a Multimode AFM system (Veeco, USA). For the experiments we used MLCT cantilevers with 4 triangular tips with different spring constants.

Figure 1: Home-made AFM system

Neutrophil polymorphonuclear granulocytes (PMNs) were isolated from blood using a standard density gradient separation method.

Soft polyamide gels of different stiffness were also prepared from acrylamide / bisacrylamide solutions and were used as reference materials.

Calibration of photodiode sensitivity

After initial laser beam alignment settings the deflection of the cantilever is monitored as a function of sample/piezo position while the tip comes in contact with the sample surface (usually a hard surface like glass). After the analysis of the data curve the measured cantilever deflection z can be converted to the corresponding force F if the spring constant of the lever k is known.

Figure 2: Determination of sensitivity and curve fitting (2)

Attention must be paid to the fact that depending on the reproducibility of the curves obtained one should repeat the procedure several times to gain accuracy. Also curve fitting and especially the correct choice of “zero” contact point may play an important role. There is an apparent to provide standardization rules to solve these problems, however this needs additional work and communication between several working groups.

Accurate determination of spring constant

The accurate determination of the spring constant is also essential in order to correctly evaluate force measurements. The cantilever’s bending stiffness or spring constant, k, is usually estimated by its manufacturer, but this estimate can be quite gross, and sometimes the error can exceed even 100%. For this reason, several methods have been developed to arrive at a more accurate estimate. In our case k was determined using the thermal tune method where the voltage noise recorded is related to the thermal fluctuations of the cantilever position and angle.

Figure 3: Frequency vs. amplitude plot from a thermal tune (taken in air).

A) full-scale spectrum. B) zoomed thermal tab  and fitted (blue line).

(figure taken from Asylum MFP 3D procedural operation manualette)

Unfortunately the result proved to be instrument-dependant for the two different types of AFM systems used, which means this is also one field subjected to the need of standardization rules.

Results and discussion

Results

Evaluation of the elastic modulus after recording the load-depth curve is also an interesting and sometimes questionable procedure, as quite a few models exist and in many cases of commercially available systems the built in software provides an output value without any further information.

In our case the elastic modulus (E) of neutrophils was in the range of a few kPas and was determined analyzing the data curve using a home-made software. The cell experiments were performed just a few hours after venipuncture, as time plays an important role in the evaluation of cell mechanical properties. In this case also, there is a need to provide rules for several parameters as fe the buffers and solutions used, fixing or labeling procedures that may be used in the case that AFM is combined with other microscopic techniques (as epifluorescent / confocal microscopy or scanning electron microscopy (SEM)), as well as for the suggested time of measurement, given that cell activation may influence cell surface mechanical properties.

On the other hand the evaluation of the gels measured yielded elastic moduli between 10 and 500 kPa. In this case the procedure was much more standardized and is suggested to be used as reference to compare measurements in the case of different labs using different instrumentation.

Discussion

Atomic Force Microscopy has been proven to be a valuable tool in the case of cell biology, not only for cell imaging, but also for evaluating cell mechanical properties, providing information on the structure and function of the cells as well as of the underlying cytoskeleton and cell organelles.

However, as it was pointed also by the outcome of this STSM, there are also a lot of parameters that can influence the results of this type of studies, making it really difficult in many cases to compare results from different groups using different AFM setups. For example tip selection and selection of mode (contact, non contact, tapping) can greatly influence the quality of the images obtained.

In the case of force measurements we may overall state that in order to get reliable results there is an urgent need to set standard methods and rules followed by the different groups in the evaluation of cell nanomechanics. Sensitivity, spring constant, data processing and evaluation of the elastic modulus are some of the important parameters we have to find standard and generally accepted ways to determine.

In addition there is also a need to get a better insight in the built-in software of the different AFM systems used (especially commercial ones) in order to be also able to compare their results.

References

1. T Ohashi, Y Ishii, Y Ishikawa, T Matsumoto, M Sato. Experimental and numerical analyses of local mechanical properties measured by atomic force microscopy for sheared endothelial cells. Biomed Mater Eng. 2002; 12 (3):319-27

2. Małgorzata Lekka. The AFM detection of ligand-receptor interaction on a surface of living cells (internet presentation)

3. Fabien Gaboriaud, Yves F. Dufrêne.  Atomic force microscopy of microbial cells: Application to nanomechanical properties, surface forces and molecular recognition forces Colloids and Surfaces B: Biointerfaces Volume 54, Issue 1, 15 January 2007, Pages 10–19

4. Tatyana G. Kuznetsova, Maria N. Starodubtseva, Nicolai I. Yegorenkov, Sergey A. Chizhik, Renat I. Zhdanov. Atomic force microscopy probing of cell elasticity Micron Volume 38, Issue 8, December 2007, Pages 824–833 Microscopy in Nanobiotechnology

5. S. Kasas, G. Dietler. Probing nanomechanical properties from biomolecules to living cells Pflügers Archiv - European Journal of Physiology April 2008, Volume 456, Issue 1, pp 13-27

6. Marija Plodinec, Marko Loparic,Rosmarie Suetterlin, Harald Herrmann,Ueli Aebi, Cora-Ann Schoenenberger. The nanomechanical properties of rat fibroblasts are modulated by interfering with the vimentin intermediate filament system Journal of Structural Biology Volume 174, Issue 3, June 2011, Pages 476–484

7. Girasole M, Dinarelli S, Boumis G. Structural, morphological and nanomechanical characterisation of intermediate states in the ageing of erythrocytes. J Mol Recognit. 2012 May;25(5):285-91.

8. Dimitrios Kirmizis, Stergios Logothetidis. Atomic force microscopy probing in the measurement of cell mechanics Int J Nanomedicine. 2010 ;5 :137-45

9. Lewis W. Francis, Paul D. Lewis, Chris J. Wright and R. Steve Conlan. Atomic force microscopy comes of age Biology of the Cell (2010) 102, (133–143)

10. Gibson, C. T., Watson, G. S. Scanning force microscopy - Calibrative procedures for 'best practice'. Scanning 19 (8): 564-581, 1997.

# 11. Palacio, M. L. B., Bhushan B. Normal and Lateral Force Calibration Techniques for AFM Cantilevers. Critical Reviews in Solid State and Materials Sciences 35 (2): 73-104, 2010.